## Slide 1: How to use Multimeter
www.discom.de

## Slide 2: What is Multimeter
Multimeter is a general-purpose measurement application with only fixed frequency analysis (no order spectra).
Sometimes you need a tool for a spontaneous measurement, where in the first place you want to do wave recordings which can later be fed into a proper analysis. Multimeter serves as such a tool.
Since it has a life display of the sensor signals and spectra, it can also be used for on-the-fly analysis, but everything interesting should be recorded.
Obviously, you need a Tas box and at least one sensor to do something useful with Multimeter.

## Slide 3: Multimeter Features
Multimeter is prepared for four sensors, one analog value (like torque) and one speed pulse channel. (Obviously, you switch on only those you need.)
The sensors are labelled VS1, VS2, Mic1 and Mic2, but this has no influence on the recording – the data go directly into the wave file, whatever they are.
In Multimeter, additional high pass filtered versions of the VS signals are generated, and A filter is applied to the Mic sensors, but the wave recordings contain the unfiltered signals.
The “analog value” channel is meant to be connected to some voltage source, for example a torque cell or a position sensor. In a normal measurement project, these “analog values” serve as control values.
The speed channel is equipped with a speed pulse decoder to retrieve a rotational speed. This speed is not used for anything (no order spectra), but if your test objects generates a pulsed speed signal (with not too many pulses per revolution), you can use the speed channel to check whether the pulse train is usable for later applications.

## Slide 4: Setup 1: Hardware
Since you want to do measurements, you need a Tas box. You can use any type of Tas Box; the settings have to be adapted anyway.
Connect the Tas box via USB to your computer, and connect the sensors to the Tas Box (any channels you like). If you have more than 4 channels, you may need to use the external 12V power adapter for the Tas box.
12V power adapter (if necessary)
USB
BNC
Sensors
All Tas boxes have rubber feet to isolate them from whatever they stand on. If you are using a notebook computer, this will in most cases also be isolated from the surrounding. This way your system is potential-free, which might or might not be helpful. If you are using the power adapter for the Tas box, this will provide a grounding.
Your Computer

## Slide 5: Setup 2: Load Parameter Database
Most of the settings you will have to do during the setup phase are somehow linked to information which comes from the parameter database.
For example, the names of all sensor signals are defined in the parameter database, so to be able to assign sensors to input channels, you need to load the parameters.
To load the parameter database into TasAlyser, you have to start a test run. More on test runs later; for now just start TasAlyser and press F5 and [OK], or press the [Inserted] button in Command Center:
This window appears on F5. Just press OK here. More details later. If the window does not appear on F5, first click on TasAlysers main window title bar, then try again.
This is the Command Center window. Press [Inserted] here.
You will probably get Tas box error messages on first try. That’s OK, just ignore them and continue on next page.

## Slide 6: Setup 3: Tas Box Settings in TasAlyser
In the Tas box settings, you specify which sensor is connected to which Tas box channel.
In the Discom system, the available signals (like “Vibration Sensor 1”, “Microphone 2”) are set up in the parameter database and used throughout the system. The Tas box settings is the only place where the real, physical inputs are linked to the signals.
Switch on those Tas box channels you want to use and select the according signal source
Select ICP coupling for accelerometers and microphones and DC for the “Analog Value”
Use an appropriate sensitivity range.
For accelerometers, 2V or 5V will work in most cases; microphones need higher sensitivity (= smaller range).
If the list in the Tas box settings is empty: see previous page.
Don’t forget to press the ‘Save’ button in the main toolbar after making changes!
Expand the FAVORITES window and double-click on TAS BOX.

## Slide 7: About Calibration
Depending on the situation, you can do a proper calibration before your measurements, enter the calibration factor(s) manually, or take care of that after the first recording.
The Calibration software module can also be found in the Favorites. Refer to the separate calibration documentation on how to perform the calibration or to enter calibration factors manually (see page 4 of “About Discom Calibration”).
Remember: we are talking about production testing. For the results, the exact absolute levels of the measured vibrations are not so important, but the repeatability and the correlation with quality (higher levels correspond to louder noise) are.
You should also adjust the channel sensitivity (gain) in the Tas box settings (see previous page), using the signal monitor or a first wave recording. If the wave recording is clipped at high signal levels, use a larger sensitivity range.
Find the Signal Monitor in the Favorites
Check the gain display for the vibration sensors and microphones.
You can also route the sensor signals to the PC sound output and listen using headphones.
Channel selection for Audio monitoring

## Slide 8: Quick Sensor Test
The Calibration module can also be used for a quick test whether the sensors are connected correctly and provide signals.
Manually start a test run (see “Step 2: Load the parameter database”) and open Calibration control from the Favorites. Then…
(1) Switch on Manual channel selection
(3) Select the channel you want to check
(2) Press Channel Info
This will open the next window.
The Channel Information window has a time signal and a spectrum display. Get someone to (lightly) hit the sensor and check whether the time signal reacts.

## Slide 9: Step 4: Check Wave Recorder Settings
Just to avoid simple mistakes: check the recorder settings and make sure that recording is activated at all!
Open Wave Recorder from the Favorites
Make sure that Recording button is Active (red)
In the Files section you can set up the recording location, file name construction and more
The Options should be set up like this for Multimeter
This check mark should only be set for manually controlled measurements (like here in Multimeter), not in automatic test runs (like in test stands)!

## Slide 10: About Test Runs
The Discom system is inherently a production test system. A measurement is called a Test Run, each test run is divided into Test Steps (also called Modes), and within each test step there is a certain time interval where a Measurement is performed and data are evaluated against limits. Each test run generates a wave file.
Also Mulitmeter uses this concepts. Since there is no test stand control (or maybe no test stand at all), you, the operator, have to initiate test runs, select modes and start measurements manually. In Multimeter, there are only two test steps, “Runin” and “NoiseTest”.
Signals and spectra are only calculated when a mode is selected.
The wave file recording happens only while the “measurement” is running.
Test Run
Mode “Runin”
Mode “NoiseTest”
Wave Recording
Read more about test run control on the next page.
“Insert”
= Start
“Remove”
= End

## Slide 11: Test Run Control
TasAlyser has several ways to control the test run. For manual operation, the keyboard shortcuts and the Command Center window are most suited.
You can use both methods in parallel, whatever suits you best!
Keyboard interface:
Press F5 to start a test run.
(If the “Start Test Run” window does not appear, click into TasAlyser’s main window title bar and try again.)
You can enter a serial number for your test run and a comment to be stored with the data.
F2 re-opens this window during a test run.
Press F6 to select a test step and F7 or <space> to start/stop measurement and wave recording.
F8 completes the test run. Second opportunity for entering a serial number.
F9 will cancel (discard) the current test run.
Command Center Interface:
Press [Inserted] to start a test run, and press it again to end the test run.
Select a test step from the list to activate it.
Press the [Measuring] button to record wave data, and again to end recording.

## Slide 12: Try your first recording!
Let’s get going!
Start a test run (F5 / [Inserted] — see previous page)
Select mode “Runin” and check whether you get any sensible signals (see next page)
Select mode “NoiseTest”, wait until the interesting noise is about to start
Start a measurement (<space> / [Measuring] — see previous page)
Wait until the interesting noise is finished, then end the measurement
End the test run (F8 / [Inserted] off — see previous page)
Press F12 to open the Wave Recorder, press the “load most recent recording” button and then the “open in TasWavEditor” button.
Wave files are created in the sub-folder ‘Recordings’ of the Multimeter project folder.
Read more about TasWavEditor in the according documentation.

## Slide 13: Multimeter Main Windows
TasAlyser has lots of different displays and windows built in and hidden somewhere. Some will be available in all projects, others only if configured in the project.
For Multimeter, this is the standard layout (which you can change as you need):
Favorites and System Configuration windows are docked at the left border.
The System Configuration gives access to all settings and all functions of TasAlyser.
Time Signals and FFT Spectra
For the two accel. sensors; mics go to panes 3 and 4.
Spectrograms are recorded only while ‘Measurement’ is running.
Read more about these controls on the next page

## Slide 14: Using the ‘Scope’
General options popup menu:
Right-click into the area outside of the dark blue display panes to call up the popup menu for general scope options.
There is also a popup menu within each pane.
y axis zoom controls
x axis zoom controls
global Autoscale button
There is a lot of functionality hidden in the scope…
Zoom controls:
zoom in
zoom out
autoscale axis
scroll to zero position
Hit me:
You can navigate around by using the keyboard arrow keys and the mouse wheel.
Use the keyboard keys + and – for zooming y axis in and out and ‘home’ key for autoscale.
More Functions:
Right-click into a pane to get a popup menu with additional functions to be applied to this pane, for example graphics export.
Data Sources:
Use the Data Sources button to assign sensor signals to display panes and to change curve colors.

## Slide 15: Signal Processing Chain (simplified)
Time Signal
The processing of sensor signals in TasAlyser works about like this (simplified, obviously):
TAS Box does A/D conversion and transmits data via USB.
USB data stream is converted to floating point data; calibration is applied
Resampling generates processing channels for fixed frequencies and rotational synchronous analysis
(Rotationally synchronous analysis is not configured into Mulitmeter.)
Fourier Transform
RMS calculation
Logarithm
Spectral Averaging
Display curves in Scope windows
Peak hold maximization, compare to limit
Sampling rate and block size parameters are applied during resampling. These are set in parameter database. (See further on in this documentation.)
Windowing parameter is applied in FFT step and set up directly in TasAlyser.
Averaging uses a running exponential average.
In TasAlyser you can see the sensor signals and the frequency spectra.
Store in rdt file

## Slide 16: Start playing!
Go ahead and play with Multimeter!
Do more recordings, try different settings, play back the wave files.
If you mess up something, close TasAlyser and restart it: no changes will be saved unless you press the ‘Save’ button.

## Slide 17: Where to go from here
If your main purpose is to view and record time signals, you are done.
If you also want to work with an analog value or a speed pulse signal, read on and look into the next pages.
After that you can find information about FFT parametrization and a short explanation on how to set up signals in the parameter data base.

## Slide 18: Setting up Analog Value Measurement
To move the Instrument, just “grab” it anywhere with the mouse.
Double-Click inside the “Instrument” to open the signal source settings.
Right-Click inside the “Instrument” to open the display settings.
Look for the “Analog Value” in the Favorites.
Setting up the analog value input (e.g. for torque) is simple: switch on the according Tas box channel (see page 3), enter an appropriate calibration factor, open the “Instrument” from the Favorites:
The “Sensor Signal” selection is linked to the signals defined in the parameter database – see next page!

## Slide 19: Setting up Speed Detection
The first steps are the same as for the Analog Value: switch on the Tas box channel, open the instrument and double-click it to access the settings.
Enter the number of pulses per revolution of the speed encoder here.
Press Pulse detection settings to access the detailed settings explained on the next page.
Speed pulse detection uses a Schmitt trigger method: there are two thresholds. A speed pulse is marked when the signal passes the lower threshold from below, but before the next pulse is accepted, the signal has to pass the upper threshold from above:
First pulse detected here
Ready for next pulse after here
Second pulse detected here
This method is robust against noise and signal fluctuations.

## Slide 20: Setting up Speed Pulse Detection
The most important setting is obviously the choice of the two trigger thresholds.
Use the [Signal Levels] and [Show last input signal] functions to get an idea about the actual speed pulse voltages and choose appropriate levels:
Enter the lower and upper Schmitt trigger level here.
Press Signal levels? while the test object is rotating to get information about the actual input.
Press Show last input signal while the test object is rotating to see a snapshot of the speed pulse signal.
The actual input curve is purple, the speed pulses green. The lower level of the green curve shows the lower trigger threshold, the peaks go to the upper threshold.

## Slide 21: FFT Parameters
The main parameters are Block Length (how many samples per block) and Sampling Frequency. They are set in the parameter database.
The base sampling rate is set in the Tas Box, but the Resampling unit can generate any other sampling rate below that value.
The combination of block length and sampling frequency determines the spectral resolution and the length of the spectra.
Press Sampling Parameters in the Main Form.
Block Length is in samples.
The spectra will have half the length.
Block Increment is also in samples and defines the overlap.
Fix Frequency is in kHz. The highest frequency in the spectrum is half the Fix Frequency.
Ignore this for now.
These settings produce spectra going to 16 kHz with 1024 points (spectral resolution 15.6 Hz).
Operation hint: all changes apply immediately with the next test run. You do not need to press a save button or close TasForms.

## Slide 22: More on FFT parameters
FFT windowing is specified in TasAlyser. Open the System Configuration and double-click on the FFT module:
Raw data stream with 50 kHz
Resampled data stream with 32 kHz
Blocks of 2048 samples, shifted by 512 samples
2048
Using the example settings from the previous page, the processing steps are like this:
Resampling
FFT produces spectra of 1024 points reaching to 16 kHz.
Advance is 512 samples corresponding to 16 ms at 32 kHz.
16 ms

## Slide 23: Signal Setup in Parameter Database
The Discom parameter database is based on Microsoft Access (due to historical and practical reasons).
There are actually two mdb database file: one contains the tables with data, the other holds the user interface. This latter one is called “TasForms”, so when you intend to edit parameters, double-click on TasForms – or the shortcut called “Parameter Database”:
The first thing you see is the Main Form with lots of buttons.
Check Show Advanced Settings and then press Signal Settings.
In the Signal Settings form press Input Signals.
This table shows all signals which are defined in this project.
You can edit properties like selecting a different unit or change the logarithmic reference.
You can NOT change the names here.
Continued on next page...
Operation hint: all changes apply immediately with the next test run. You do not need to press a save button or close TasForms.

## Slide 24: Changing Signal Definitions
The standard setup has pre-configured sensors called “VS1”, VS2”…, “Mic1” and “Mic2”,
suggesting vibration sensors and microphones, but you can use these sensors in any way you need. Change the properties (like Unit) in the Signal Properties form and change the name directly in the database table:
“Unit” consists of the base unit (like m/s² or Pa) and a prefix like µ, m, k. ‘1’ is no prefix.
Example:
for m/s², select prefix ‘1’.
for mm/s², select prefix ‘m’.
Use “Copy only basic information” if you only need a new signal name.
With “Copy all connected data”, all measurement definitions which exists for the “Copy from” sensor will also be copied.
A new signal definition is added by copying an existing signal and then modifying it as necessary.
Open table ApplicationSignal directly in Access (press F11 in TasForms).
Change the names here.

## Slide 25: Signal Groups
“Noise” sensors (which you want to get spectra from) are organized in Signal Groups.
To get the data of a “noise” sensor processed, you have to put it into a signal group.
The Resampling is parametrized for signal groups.
If you have created a new sensor, add a line to the signal grouping list: select a signal group in the first column and the new sensor in the last. This places the new sensor in that group.
Sensors can appear in more than one signal group.
Sensor VS3 is in signal group ‘Vibration Sensors’.

## Slide 26: Finish
That’s all about Multimeter.
Obviously, there is a lot more to learn and to play with, but that’s another story for another time.
Play with the settings in multimeter and with the application itself to get a feeling for how it is operated. Use more sensors. Work with TasWavEditor. Have fun!
The next project to explore could be the Demo-NickSynchResamp project, which shows the effect of rotationally synchronous resampling, and afterwards the Demo-EDrive project which explains more about this topic and how to set up the parameters.